## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
1BR-IC | 1 Bed Room - Intensive Care | 25 | 24 Hours | 1 patient, 1 visitor, up to 6-8 staff |  | Acoustic privacy is required | An acuity adaptable, enclosed bed room for a patient requiring intensive medical treatment, nursing care and monitoring for potentially life threatening conditions. Visual access to a main Staff Station or Sub Staff Station is required. Natural light is required to patient beds. Dialysis facilities may be required, assumed to be provided via mobile reverse osmosis water dialysis units. As mobile Imaging may be required, shielding requirements must be confirmed with Radiation Consultant.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
1BR-IC | 1 Bed Room - Intensive Care | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
1BR-IC | 1 Bed Room - Intensive Care | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
1BR-IC | 1 Bed Room - Intensive Care | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
1BR-IC | 1 Bed Room - Intensive Care | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
1BR-IC | 1 Bed Room - Intensive Care | FA | WLPR-012 | WALL PROTECTION: corner guards, to 900 AFFL, prefinished PVC | 1
1BR-IC | 1 Bed Room - Intensive Care | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
1BR-IC | 1 Bed Room - Intensive Care | FA | DWWH-005 | WINDOW | 1
1BR-IC | 1 Bed Room - Intensive Care | FA | DWGL-002 | DOOR: 910mm c/o, single leaf, half glazed, solid core/glass, paint | 1 | optional, additional door to adjoining space
1BR-IC | 1 Bed Room - Intensive Care | FA | DWSC-012 | DOOR: 1200mm c/o, 1 1/2 leaf, solid core, paint | 1 | optional to ensuite
1BR-IC | 1 Bed Room - Intensive Care | FA | DWGL-071 | DOOR: triple leaf, sliding, fully glazed, alum/safety glass, powdercoat/paint | 1
1BR-IC | 1 Bed Room - Intensive Care | FA | DWOB-006 | OBSERVATION: internal window, alum/safety glass, double glazed, integral venetian blind, powdercoat | 1 | to partition walls
1BR-IC | 1 Bed Room - Intensive Care | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
1BR-IC | 1 Bed Room - Intensive Care | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 11
1BR-IC | 1 Bed Room - Intensive Care | FFE | MGFP-002 | FLOWMETER: medical air | 3 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | DWWD-007 | BLIND: venetian, integral | 4 | to window and observation panel between rooms and to sub staff station, controls to both sides of internal observation windows
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIHR-019 | BRACKET: for suction bottle | 4
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1 | for staff
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-019 | CHAIR: patient | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQCW-035 | BIN: waste, recycling, plastics, 50L | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIRT-031 | TRACK: curtain, privacy | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQGE-010 | CURTAIN: privacy screen | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQCW-016 | BIN: waste, general, 20L | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQCW-004 | BIN: waste, clinical, 120L | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near entry
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMOP-027 | IV POLE: on pendant | 1 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQTA-013 | TABLE: overbed | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MGFP-004 | FLOWMETER: oxygen | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-142 | TROLLEY: sharps bin | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 3 | wire basket bins
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-002 | ADAPTER: suction | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-100 | SUCTION BOTTLE | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMBE-013 | BED: patient, ICU, electric | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | optional
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMSP-020 | PENDANT: tandem, wet and dry combination (fixed or adjustable height) | 1 | with drawer for patient property
1BR-IC | 1 Bed Room - Intensive Care | FFE | ITCL-014 | CLOCK: analogue, electric | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-049 | MACHINE: dialysis | 1 | optional, if mobile dialysis required
1BR-IC | 1 Bed Room - Intensive Care | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMHA-023 | VENTILATOR | 1 | on pendant 1-A
1BR-IC | 1 Bed Room - Intensive Care | FFE | ITBU-006 | CONTROL PANEL: patient entertainment, handset | 1 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | optional
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMHA-030 | MONITOR: patient, cardiac, pendant mounted | 1 | on pendant 1-A
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQBS-055 | CHAIR: overnight fold out | 1 | can optionally be a visitor chair
1BR-IC | 1 Bed Room - Intensive Care | FFE | FIJO-213 | PANEL: joinery carcass end | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | FQSM-043 | TROLLEY: nested, set of 3, underbench | 1
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMBE-068 | LIFTER: patient, pendant mounted | 1 | optional
1BR-IC | 1 Bed Room - Intensive Care | FFE | MMGE-178 | PROCESSOR: reverse osmosis RO water, mobile | 1 | optional, if mobile dialysis required
1BR-IC | 1 Bed Room - Intensive Care | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELPR-006 | RCD: residual current device | 2 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | for display screen
1BR-IC | 1 Bed Room - Intensive Care | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on tandem pendant, + indicator button & light
1BR-IC | 1 Bed Room - Intensive Care | SE | HYTP-008 | OUTLET: cold water | 1 | optional, for reverse osmosis water unit
1BR-IC | 1 Bed Room - Intensive Care | SE | LIFX-017 | LIGHT: night light | 2 | wall washers
1BR-IC | 1 Bed Room - Intensive Care | SE | ELBO-015 | SWITCH: light, single | 3 | optional, to ensuite and anteroom if required, one to night light
1BR-IC | 1 Bed Room - Intensive Care | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on tandem pendant, + indicator button & light
1BR-IC | 1 Bed Room - Intensive Care | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on tandem pendant, call button on the handset
1BR-IC | 1 Bed Room - Intensive Care | SE | ITCL-004 | BUTTON: nurse call, emergency | 2 | on tandem pendant, + indicator button & light
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-109 | GPO: single, wall mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4 | optional, if mobile dialysis required
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELPR-004 | MONITOR: line isolation | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELBO-012 | SWITCH: light, dimmer | 2 | one on pendant, one to wall
1BR-IC | 1 Bed Room - Intensive Care | SE | ITIN-014 | OUTLET: data, single RJ45, on pendant | 6 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | ITIN-071 | OUTLET: MATV, ceiling mounted | 1 | optionally can be a data outlet, for patient entertainment system
1BR-IC | 1 Bed Room - Intensive Care | SE | MGAS-021 | OUTLET: medical air (MA), on pendant | 3 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | MGAS-041 | OUTLET: oxygen (O2), on pendant | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | MGAS-061 | OUTLET: suction, on pendant | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-132 | GPO: UPS, single, on pendant | 4 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-121 | GPO: emergency power, single, on pendant | 16 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
1BR-IC | 1 Bed Room - Intensive Care | SE | ELPR-007 | UPS: status indicator | 2 | on tandem pendant
1BR-IC | 1 Bed Room - Intensive Care | SE | HYDR-015 | DRAIN: tundish unit, integral drainage outlets | 1 | optional, if mobile dialysis required
1BR-IC | 1 Bed Room - Intensive Care | SE | HYDR-016 | DRAIN: floor waste, sealed | 1