## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
ANCRT | Angiography Control/ Reporting Room | 14 | Available 24 Hours | Up to 3 staff |  | Acoustic privacy is required | Room adjacent to Angiography Procedure Room/s with direct visibility to the patient for remote control of equipment, and review and reporting of procedure images. The Control/ Reporting room may serve two procedure rooms. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts from Angiography Procedure Room and Computer/Equipment room may be required.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
ANCRT | Angiography Control/ Reporting Room | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to procedure room | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | DWPR-003 | DOOR PROTECTION: kickplate, to 300 AFFL, vinyl, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | ELBO-015 | SWITCH: light, single | 2 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | ELGP-106 | GPO: single, special, for equipment | 1 | Angiography console equipment | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | ELGP-208 | GPO: double, wall mounted | 9 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FIBM-013 | NOTICEBOARD: fixed, fabric covered | 3 | over bench | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FIHR-026 | HOOK: coat | 2 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FIJO-044 | BENCH: special depth, open under | 3 | 800/600D, at 900H, laminate finish | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 6 | for books, manuals, over bench | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 3 |  | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1 | under bench | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 6 | 3 outlets under bench for telephone | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 |  | Yes | Yes | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | ITNE-023 | PRINTER: networked | 1 |  | Yes | Yes | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2 | on bench | No | Yes | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | with handsfree microphone to Angio room | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | SE | LIFX-008 | LIGHT: task light, built in | 4 | under O/H shelving | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | MMMI-013 | CONSOLE: control, examination control console, angiography single plane ceiling | 1 | Angiography control monitors | Yes | Yes | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | MMMI-356 | DISPLAY SCREEN: dual screen, angiography | 1 |  | Yes | Yes | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FFE | WLPT-003 | PARTITION: laminate | 2 | as required | No | No | No | No | No | No | No | No | No
ANCRT | Angiography Control/ Reporting Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No