## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
ANPR | Angiography Procedure Room | 42 | Available 24 Hours | 1 patient, up to 6 staff |  | Acoustic privacy is required | Procedure room where Angiography procedures and electrophysiology studies are undertaken in a controlled environment. A Control room will be located adjacent and may serve more than one procedure room. A clinical scrub area should be located immediately adjacent to the room staff entry. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Unistrut Ceiling able to carry moving loads may be required. Cable ducts from Control Room and Computer/Equipment room may be required. 'X-Ray in Use' lights are required outside each entry door. Power to Angiography equipment should be on emergency supply.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
ANPR | Angiography Procedure Room | FA | CLCN-007 | CORNICE: plasterboard, flush coved between wall and ceiling, paint | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | at 3000 mm AFFL | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 1 | staff entry/ exit | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to control room | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | DWSC-026 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, lead shielding, paint | 1 | patient entry/ exit | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELBO-015 | SWITCH: light, single | 3 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-105 | GPO: single, on services panel | 7 | on Services Panel 1 | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-105 | GPO: single, on services panel | 1 | on Services Panel 3, to angiography table, 20 amp | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | on Services Panel 1 | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1 | on Services Panel 3, 20 amp | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1 | on Services Panel 1, 20 amp | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-222 | GPO: emergency power, double, on services panel | 1 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | 2 above bench | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | below services panel | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIJO-083 | BENCH TOP: special depth, joinery under, laminate finish | 1 | 540D | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1 | under bench | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIJO-123 | CUPBOARD: tall, double sliding door, glazed, catheter hanging | 1 | sliding glass doors, shelving to one side | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on Services Panel 3, indicator button & light, | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on Services Panel 3, indicator button & light, | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITCL-009 | BUTTON: nurse call, staff / nurse assist | 1 | on Services Panel 3, indicator button & light, | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1 |  | Yes | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | on Services Panel 1 | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1 outlet for telephone | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on Services Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | on Services Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on Services Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | on Services Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on Services Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | services panel | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | services panel | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMGE-001 | ADAPTER: scavenge | 1 | services panel | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMGE-002 | ADAPTER: suction | 2 |  | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | on bracket | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMHA-005 | MONITOR: patient, cardiac | 3 | on ceiling support frame | Yes | Yes | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMMI-024 | TABLE: operating, fixed, narrow, angiography single plane ceiling | 1 |  | Yes | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMMI-028 | SYSTEM: imaging, angiography single plane ceiling | 1 | with imaging & cardiac monitors | Yes | Yes | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MMSP-025 | SERVICES PANEL: special | 1 | Panel 2, to angiography table | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 2 | Panels 1 and 3 | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FFE | MMSP-029 | LIGHT: operating, single, 550 diam nom. | 1 |  | Yes | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | WLPR-019 | WALL PROTECTION: crash rail, at 900 AFFL, prefinished PVC | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified | No | No | No | No | No | No | No | No | No
ANPR | Angiography Procedure Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No