## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
FLUO | Fluoroscopy Room | 40 | Available 24 Hours | 1 patient; 1 - 2 staff |  | Acoustic privacy is required | The Fluoroscopy Room is a room for undertaking radiographic procedures involving contrast media. General anaesthesia may be given in this room. A control room is required accessed from the Fluoroscopy Room. An Equipment Room for equipment generating noise and heat is to be considered subject to equipment manufacturer. The Control Room and Equipment Room will need to be located in close proximity due to acceptable cable lengths. Power and control cabinets may be located within the room. An accessible toilet with changing facilities and shower accessed directly from the Fluoroscopy Room is required. A preparation area with sink is required in the room. The room may be used as an overflow general x-ray room. Additional Design Requirements: Radiation shielding to walls and lead lined glass as advised by Radiation Consultant. Shielded observation window will be required where patient observation cameras are not in use. Unistrut support at ceiling level and 3000mm AFFL ceiling may be required, depending on equipment to be installed. Cable ducting between the imaging unit table, equipment and control area will be required, depending on equipment type and layout. 'X-ray in Use' warning lights will be required outside each entry door. Equipment access route clearances to be considered.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
FLUO | Fluoroscopy Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
FLUO | Fluoroscopy Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
FLUO | Fluoroscopy Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
FLUO | Fluoroscopy Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
FLUO | Fluoroscopy Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
FLUO | Fluoroscopy Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2 | to prep. bench, basin & sink
FLUO | Fluoroscopy Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
FLUO | Fluoroscopy Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass, to control room
FLUO | Fluoroscopy Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 2
FLUO | Fluoroscopy Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
FLUO | Fluoroscopy Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
FLUO | Fluoroscopy Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 11
FLUO | Fluoroscopy Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | to services panel
FLUO | Fluoroscopy Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | wall mounted
FLUO | Fluoroscopy Room | FFE | MMGE-019 | CHAIR: patient | 1
FLUO | Fluoroscopy Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
FLUO | Fluoroscopy Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2
FLUO | Fluoroscopy Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2
FLUO | Fluoroscopy Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
FLUO | Fluoroscopy Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
FLUO | Fluoroscopy Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
FLUO | Fluoroscopy Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
FLUO | Fluoroscopy Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
FLUO | Fluoroscopy Room | FFE | MMGE-113 | TROLLEY: dressing | 1
FLUO | Fluoroscopy Room | FFE | FIHR-044 | RACK: lead apron | 1 | wall mounted
FLUO | Fluoroscopy Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
FLUO | Fluoroscopy Room | FFE | FIHR-011 | BRACKET: for monitor, wall mounted | 1
FLUO | Fluoroscopy Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | to services panel
FLUO | Fluoroscopy Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
FLUO | Fluoroscopy Room | FFE | MMGE-002 | ADAPTER: suction | 2
FLUO | Fluoroscopy Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
FLUO | Fluoroscopy Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
FLUO | Fluoroscopy Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
FLUO | Fluoroscopy Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
FLUO | Fluoroscopy Room | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
FLUO | Fluoroscopy Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
FLUO | Fluoroscopy Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
FLUO | Fluoroscopy Room | FFE | MMHA-005 | MONITOR: patient, cardiac | 1 | wall mounted, can optionally be on a trolley
FLUO | Fluoroscopy Room | FFE | HYBA-051 | SINK: single | 1 | warm water
FLUO | Fluoroscopy Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
FLUO | Fluoroscopy Room | FFE | FIJO-060 | BENCH TOP: 600D, joinery under, stainless steel lipped | 1 | wet edge
FLUO | Fluoroscopy Room | FFE | MMMI-283 | BUCKY: wall, general X-Ray | 1 | optional
FLUO | Fluoroscopy Room | FFE | MMMI-121 | INJECTOR: mobile, fluoroscopy/screening | 1 | optional, if angiography procedures are undertaken
FLUO | Fluoroscopy Room | FFE | MMMI-122 | DISPLAY SCREEN: for fluoroscopy/screening | 2 | on articulated arm, ceiling mounted, no. of screens subject to equipment setup
FLUO | Fluoroscopy Room | FFE | MMMI-123 | SHIELD: radiation, upper body, mobile, fluoroscopy/screening | 1
FLUO | Fluoroscopy Room | FFE | MMMI-289 | STAND: ceiling, with traverse bridge, general X-Ray | 1
FLUO | Fluoroscopy Room | FFE | MMMI-124 | STAND: floor, with patient table, fluoroscopy/screening | 1
FLUO | Fluoroscopy Room | FFE | MMMI-128 | CONTROL UNIT: fluoroscopy, mobile | 1
FLUO | Fluoroscopy Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
FLUO | Fluoroscopy Room | FFE | MMSP-036 | PENDANT: single, dual arm, dual display screen | 1
FLUO | Fluoroscopy Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
FLUO | Fluoroscopy Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to prep bench, one outlet for telephone
FLUO | Fluoroscopy Room | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
FLUO | Fluoroscopy Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1
FLUO | Fluoroscopy Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel
FLUO | Fluoroscopy Room | SE | ELBO-001 | BUTTON: emergency stop | 1
FLUO | Fluoroscopy Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | on services panel
FLUO | Fluoroscopy Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | number determined by number of doors, auto on when modality is in use
FLUO | Fluoroscopy Room | SE | ELBO-015 | SWITCH: light, single | 2 | to accessible change/wc/shwr.
FLUO | Fluoroscopy Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
FLUO | Fluoroscopy Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | on services panel
FLUO | Fluoroscopy Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
FLUO | Fluoroscopy Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
FLUO | Fluoroscopy Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel
FLUO | Fluoroscopy Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | clock
FLUO | Fluoroscopy Room | SE | ELGP-208 | GPO: double, wall mounted | 2 | to prep bench
FLUO | Fluoroscopy Room | SE | ELGP-122.1 | GPO: emergency power, single, on services panel (V) | 2
FLUO | Fluoroscopy Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
FLUO | Fluoroscopy Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
FLUO | Fluoroscopy Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | flush mounted
FLUO | Fluoroscopy Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified
FLUO | Fluoroscopy Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | to room
FLUO | Fluoroscopy Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
FLUO | Fluoroscopy Room | SE | ELGP-105 | GPO: single, on services panel | 4
FLUO | Fluoroscopy Room | SE | ELGP-130 | GPO: UPS, single | 1