## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
GENXR | General X-Ray Room | 38 | Available 24 Hours | 1 patient; 1 - 2 staff |  | Acoustic privacy is required | General X-ray is a room for undertaking general radiographic imaging, and may be located within the Imaging Unit or within an Emergency Unit. Additional Design Requirements: Radiation shielding to walls and glazing as advised by Radiation Consultant. Unistrut type ceiling may be required, depending on equipment to be installed. Cable ducting between the X-ray table, vertical bucky and control area will be required, depending on equipment type and layout. 'X-ray in Use' warning lights will be required outside each entry door. If located within the ED, quantity of services may need to align with Acute Treatment bays for trauma use.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
GENXR | General X-Ray Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
GENXR | General X-Ray Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 14
GENXR | General X-Ray Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 4 | optional
GENXR | General X-Ray Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
GENXR | General X-Ray Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 6
GENXR | General X-Ray Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin; selection of splashback material to comply with AHFG Part D Infection Control
GENXR | General X-Ray Room | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
GENXR | General X-Ray Room | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1
GENXR | General X-Ray Room | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 1
GENXR | General X-Ray Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
GENXR | General X-Ray Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
GENXR | General X-Ray Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 12
GENXR | General X-Ray Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
GENXR | General X-Ray Room | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
GENXR | General X-Ray Room | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
GENXR | General X-Ray Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
GENXR | General X-Ray Room | FFE | MMMI-350 | SCREEN: radiation, lead glass | 1 | safety glass
GENXR | General X-Ray Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
GENXR | General X-Ray Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
GENXR | General X-Ray Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
GENXR | General X-Ray Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
GENXR | General X-Ray Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
GENXR | General X-Ray Room | FFE | FIHR-044 | RACK: lead apron | 1
GENXR | General X-Ray Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
GENXR | General X-Ray Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1 | to services panel
GENXR | General X-Ray Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
GENXR | General X-Ray Room | FFE | MMGE-002 | ADAPTER: suction | 1 | to services panel
GENXR | General X-Ray Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
GENXR | General X-Ray Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
GENXR | General X-Ray Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
GENXR | General X-Ray Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
GENXR | General X-Ray Room | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 1
GENXR | General X-Ray Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
GENXR | General X-Ray Room | FFE | MMBE-010 | BED: patient, electric with mattress | 1
GENXR | General X-Ray Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
GENXR | General X-Ray Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1
GENXR | General X-Ray Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 | to Control
GENXR | General X-Ray Room | FFE | MMMI-285 | CONSOLE: control, general X-Ray | 1
GENXR | General X-Ray Room | FFE | MMMI-293 | TABLE: floating top, general X-Ray | 1 | refer to manufacturer requirements, can optionally be elevating table
GENXR | General X-Ray Room | FFE | MMMI-283 | BUCKY: wall, general X-Ray | 1
GENXR | General X-Ray Room | FFE | MMMI-287 | GENERATOR: radiology, general X-Ray | 1
GENXR | General X-Ray Room | FFE | MMMI-289 | STAND: ceiling, with traverse bridge, general X-Ray | 1
GENXR | General X-Ray Room | FFE | MMMI-295 | SCREEN: imaging stitching barrier | 1 | optional, may include one per facility
GENXR | General X-Ray Room | FFE | FIJO-029 | BENCH: 750D, open under, X-ray | 1
GENXR | General X-Ray Room | FFE | FIJO-099 | CAC: cable access cap | 2
GENXR | General X-Ray Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
GENXR | General X-Ray Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | to Control
GENXR | General X-Ray Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
GENXR | General X-Ray Room | FFE | FIHR-076 | BRACKET: for display screen, triple, desk mounted | 1
GENXR | General X-Ray Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
GENXR | General X-Ray Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 2
GENXR | General X-Ray Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | for phone
GENXR | General X-Ray Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | to radiation consultant's requirements
GENXR | General X-Ray Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1
GENXR | General X-Ray Room | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
GENXR | General X-Ray Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
GENXR | General X-Ray Room | SE | ELGP-135 | GPO: UPS, single, wall mounted | 3
GENXR | General X-Ray Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 3 | number determined by number of doors, auto on when modality is in use
GENXR | General X-Ray Room | SE | ELBO-015 | SWITCH: light, single | 1
GENXR | General X-Ray Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | on services panel; if located within the ED, quantity of services may need to align with acute treatment bays for trauma use
GENXR | General X-Ray Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | on services panel; if located within the ED, quantity of services may need to align with acute treatment bays for trauma use
GENXR | General X-Ray Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 2
GENXR | General X-Ray Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | for clock
GENXR | General X-Ray Room | SE | ELGP-208 | GPO: double, wall mounted | 2 | to control
GENXR | General X-Ray Room | SE | ELGP-122.1 | GPO: emergency power, single, on services panel (V) | 3
GENXR | General X-Ray Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
GENXR | General X-Ray Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
GENXR | General X-Ray Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
GENXR | General X-Ray Room | SE | ITSE-008 | BUTTON: duress, mobile device | 1
GENXR | General X-Ray Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
GENXR | General X-Ray Room | SE | WLSH-004 | SHIELDING: lead lining | 1
GENXR | General X-Ray Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | optional
GENXR | General X-Ray Room | SE | ELGP-105 | GPO: single, on services panel | 4 | on services panel
GENXR | General X-Ray Room | SE | ELGP-140 | GPO: three phase, single | 1