## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PBTR-R | Patient Bay, Emergency - Resuscitation | 25 | 24 Hours | 1 patient; 4 - 5 staff |  |  | A Patient Bed Bay for assessment, resuscitation and treatment of patients with critical conditions, located within the Emergency Unit. The Bay may have partition side walls and a bench used for preparation and write-up. Resuscitation and diagnostic equipment should be located in close proximity which may include ventilator, cardiac pacemaking equipment, patient sliders and mobile X-ray machine. Accurate control of ambient temperature is required. Lead shielding may be required depending on use of X-ray equipment; to be assessed by Radiation Consultant.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | WLPR-019 | WALL PROTECTION: crash rail, at 900 AFFL, prefinished PVC | 1 | as required
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | WLPR-013 | WALL PROTECTION: corner guards, to 1500 AFFL, prefinished PVC | 3 | as required
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 5
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MGFP-002 | FLOWMETER: medical air | 1 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQGE-010 | CURTAIN: privacy screen | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQCW-016 | BIN: waste, general, 20L | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1 | optional, for slave display screen
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-014 | DISPENSER: disposable glove | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIDI-021 | DISPENSER: mask | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | optional, for slave display screen
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-126 | TROLLEY: intubation | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-039 | HANGER: IV, for ceiling track | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIRT-020 | TRACK: ceiling mounted, IV | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MGFP-004 | FLOWMETER: oxygen | 3 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMHA-027 | TROLLEY: resuscitation, adult | 1 | where more than one resuscitation bay may be centrally located
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1 | optional, where more than one resuscitation bay may be centrally located
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQBS-052 | STOOL: mobile, high, adj height | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-002 | ADAPTER: suction | 3 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-100 | SUCTION BOTTLE | 3 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQSM-039 | TROLLEY: waste bag | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 1 | for patient management system
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1 | refer BHWS-A (similar)
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMBE-011 | BED: patient, emergency, electric | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITCL-012 | CAMERA: observation | 2 | optional, for telemedecine
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMHA-002 | DEFIBRILLATOR | 1 | on resuscitation trolley
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 2 | optional, one to wall, one to ceiling
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-036 | FRIDGE: drugs, underbench | 1 | optional, where more than one resuscitation bay may be centrally located
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMSP-018 | PENDANT: single, wet and dry combination (fixed or adjustable height) | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | to bench
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMHA-023 | VENTILATOR | 1 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1 | open under part
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITCL-020 | CLOCK: clock system slave, dual digital / elapsed time | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIJO-099 | CAC: cable access cap | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMHA-030 | MONITOR: patient, cardiac, pendant mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIJO-216 | CUPBOARD: pigeon holes, under bench | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FQSM-043 | TROLLEY: nested, set of 3, underbench | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIHR-076 | BRACKET: for display screen, triple, desk mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-175 | DIAGNOSTIC SET: pendant mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-176 | PUMP: infusion, pendant mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMGE-177 | PUMP: syringe, pendant mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | FFE | MMSP-037 | LIGHT: operating, single, 550 diam nom. pendant mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 4 | optional, where supplying data to drug fridge and slave display screens
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELPR-006 | RCD: residual current device | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on pendant + indicator button & light
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELBO-015 | SWITCH: light, single | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on pendant + indicator button & light
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button on the handset
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on pendant + indicator button & light
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5 | one to resus. trolley
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 1 | optional, for optional drug fridge
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 5
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITCL-023 | CONTROL PANEL: master clock system | 1
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ITIN-034 | OUTLET: data, triple RJ45, on pendant | 2
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | MGAS-021 | OUTLET: medical air (MA), on pendant | 1 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | MGAS-031 | OUTLET: nitrous oxide (N2O), on pendant | 1 | optional, on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | MGAS-041 | OUTLET: oxygen (O2), on pendant | 3 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | MGAS-061 | OUTLET: suction, on pendant | 3 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | MGAS-051 | OUTLET: scavenge, on pendant | 1 | optional, on pendant, if N2O provided
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-132 | GPO: UPS, single, on pendant | 4
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELGP-121 | GPO: emergency power, single, on pendant | 8 | on pendant
PBTR-R | Patient Bay, Emergency - Resuscitation | SE | ELPR-007 | UPS: status indicator | 1