## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PROC-20 | Procedure Room, 20m2 | 20 | 24 Hours | 1 patient; 2-4 staff |  | Sound attenuation level - High. Acoustic privacy is required. | The Procedure Room may be required for the performance of procedures that do not require the full facilities of the Operating Suite. This may include suturing or wounds, dressings, administration of local anaesthesia and use of medical gases in connection with the procedures being performed. General anaesthesia will not be administered in this room. Additional Design Considerations: Lead shielding may be required depending on use of mobile Imaging equipment. Treatment of bariatric patients including mobile, bariatric lifter access.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
PROC-20 | Procedure Room, 20m2 | FA | CLCN-007 | CORNICE: plasterboard, flush coved between wall and ceiling, paint | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 1 | Between above bench joinery and ceiling. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | Cover to glazing. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Optional. To be decided at project level. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | DWWD-003 | BLIND: integral | 1 | Optional, to external windows. Number to suit window configuration. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | DWWH-005 | WINDOW | 1 | optional | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELBO-015 | SWITCH: light, single | 2 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELGP-105 | GPO: single, on services panel | 3 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELGP-109 | GPO: single, wall mounted | 2 | low level for optional bed or exam couch, clock | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | for monitor | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | to bench, computer PC + spare. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-022 | DISPENSER: moisturiser | 1 | to hand basin. Depends on operational policy. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIHR-011 | BRACKET: for monitor | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIHR-019 | BRACKET: for suction bottle | 2 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1 | Optional . Length to suit layout, refer RLS or project specific requirements. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 1 | Optional, lockable. Number to suit layout. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2 | adjustable shelf, optional. Number to suit layout, refer RLS or project specific requirements. Storage solution for consumables and equipment to meet project specific Infection Control and WH&S requirements. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2 | over bench, adjustable shelf, optional. Number to suit layout, refer RLS or project specific requirements. Storage solution for consumables and equipment to meet project specific Infection Control and WH&S requirements. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FIRT-026 | TRACK: curtain, bedscreen, Z shape | 1 | to door opening | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | Under basin. Significant change in colour to be avoided. | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FQBS-052 | STOOL: mobile, high, adj height | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1 | wall mounted | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | FQGE-009 | CURTAIN: bed screen | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1 | elbow action levers or handsfree optional. | No | No | Yes | No | Yes | Yes | Yes | No | No
PROC-20 | Procedure Room, 20m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin | No | No | Yes | No | Yes | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel + indicator button & light | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | ITCL-014 | CLOCK: analogue, electric | 1 |  | Yes | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel, call button on the handset | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | for monitor communications with associated cabling and connections; verify outlet type/s with monitor selected | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | On bench or wall mounted, optional. If wall mounted add bracket. PACS images may be viewed, refer specific project requirements. | Yes | Yes | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1 | to bench or wall mounted | No | Yes | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | optional in ambulatory care area | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | LIFX-013 | LIGHT: procedure, ceiling mounted, small | 1 |  | Yes | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | on services panel, optional | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | on services panel, optional, if N20 required | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MGFP-002 | FLOWMETER: medical air | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MGFP-004 | FLOWMETER: oxygen | 2 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMBE-010 | BED: patient, electric with mattress | 1 |  | Yes | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMGE-001 | ADAPTER: scavenge | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMGE-002 | ADAPTER: suction | 2 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMGE-100 | SUCTION BOTTLE | 2 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMGE-119 | TROLLEY: equipment | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FFE | MMHA-008 | MONITOR: patient, cardiac with vital signs | 1 | optional | Yes | Yes | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1 | or similar | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
PROC-20 | Procedure Room, 20m2 | FA | WLWA-007 | WALL FINISH: splashback | 1 | Selection of material to comply with Part D Infection Prevention and Control AusHFG. | No | No | No | No | No | No | No | No | No